## MGIR250124_CB Sample DOC
Report ID | Instrument Name | Instrument Version | Schedule Name | Sample Timestamp | Sample Number | Sample ID | Sample Type | Mean (per analysis type) | Replicate Count | Vial Position | PPM | Adjusted ABS | Sample Pressure (psig) | Method Name
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:02 | 1 | Clean | Clean |  | 4 | (Clean) |  | -5.43 | 49.887
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:02 | 1 | Clean | Clean |  | 4 | (Clean) |  | 39.36 | 50.666
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:02 | 1 | Clean | Clean |  | 4 | (Clean) |  | 8.51 | 51.049
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:02 | 1 | Clean | Clean |  | 4 | (Clean) |  | 4.64 | 50.151
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:25 | 2 | Clean | Clean |  | 4 | (Clean) |  | 10.65 | 49.622
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:25 | 2 | Clean | Clean |  | 4 | (Clean) |  | 10.93 | 50.165
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:25 | 2 | Clean | Clean |  | 4 | (Clean) |  | 3.75 | 50.977
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:25 | 2 | Clean | Clean |  | 4 | (Clean) |  | 2.72 | 51.085
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:47 | 3 | Reagent/Acid Blank | Blank |  | 6 | (Blank) |  | 9.06 | 49.704
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:47 | 3 | Reagent/Acid Blank | Blank |  | 6 | (Blank) |  | 6.85 | 50.071
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:47 | 3 | Reagent/Acid Blank | Blank |  | 6 | (Blank) |  | 3.21 | 50.346
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:47 | 3 | Reagent/Acid Blank | Blank |  | 6 | (Blank) |  | 2.02 | 50.693
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:47 | 3 | Reagent/Acid Blank | Blank |  | 6 | (Blank) |  | 8.25 | 50.664
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 17:47 | 3 | Reagent/Acid Blank | Blank |  | 6 | (Blank) |  | 1.75 | 50.295
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 18:21 | 4 | Reagent/Acid Blank | Blank |  | 6 | (Blank) |  | 7.6 | 49.572
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 18:21 | 4 | Reagent/Acid Blank | Blank |  | 6 | (Blank) |  | 7.27 | 51.037
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 18:21 | 4 | Reagent/Acid Blank | Blank |  | 6 | (Blank) |  | 2.92 | 50.245
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 18:21 | 4 | Reagent/Acid Blank | Blank |  | 6 | (Blank) |  | 1.78 | 50.032
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 18:21 | 4 | Reagent/Acid Blank | Blank |  | 6 | (Blank) |  | 8.27 | 49.998
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 18:21 | 4 | Reagent/Acid Blank | Blank |  | 6 | (Blank) |  | 1.87 | 49.964
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 18:55 | 5 | Blank | Samples | 0.1115 | 3 | B | 0.082 | 5.29 | 50.6 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 18:55 | 5 | Blank | Samples | 0.1115 | 3 | B | 0.1396 | 6.16 | 50.574 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 18:55 | 5 | Blank | Samples | 0.1115 | 3 | B | 0.113 | 5.76 | 50.651 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 19:13 | 6 | MDL (0.2) | Samples | 0.274 | 2 | 1 | 0.295 | 8.5 | 50.52 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 19:13 | 6 | MDL (0.2) | Samples | 0.274 | 2 | 1 | 0.2529 | 7.86 | 50.57 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 19:27 | 7 | QCS (18) | Samples | 17.5144 | 2 | C | 17.0337 | 260.83 | 50.47 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 19:27 | 7 | QCS (18) | Samples | 17.5144 | 2 | C | 17.995 | 275.32 | 50.79 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 19:40 | 8 | QCB | Samples | 0.0595 | 2 | B | 0.0553 | 4.89 | 50.681 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 19:40 | 8 | QCB | Samples | 0.0595 | 2 | B | 0.0637 | 5.01 | 50.762 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 19:54 | 9 | CCV1 (10) | Samples | 9.7903 | 2 | A | 9.7991 | 34.63 | 50.66 | TOC Waste Water 1;10 Dilution
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 19:54 | 9 | CCV1 (10) | Samples | 9.7903 | 2 | A | 9.7815 | 34.58 | 50.809 | TOC Waste Water 1;10 Dilution
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 20:07 | 10 | CCB1 | Samples | 0 | 2 | B | 0 | 3.86 | 50.474 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 20:07 | 10 | CCB1 | Samples | 0 | 2 | B | 0 | 3.94 | 50.554 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 20:21 | 11 | MWR N/S | Samples | 54.8464 | 2 | 2 | 53.9766 | 167.82 | 50.47 | TOC Waste Water 1;10 Dilution
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 20:21 | 11 | MWR N/S | Samples | 54.8464 | 2 | 2 | 55.7162 | 173.06 | 50.785 | TOC Waste Water 1;10 Dilution
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 20:34 | 12 | CCV2 | Samples | 9.9995 | 2 | A | 9.7554 | 151.11 | 50.513 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 20:34 | 12 | CCV2 | Samples | 9.9995 | 2 | A | 10.2435 | 158.47 | 50.65 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 20:47 | 13 | CCB2 | Samples | 0.0512 | 2 | B | 0.048 | 4.77 | 50.652 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 20:47 | 13 | CCB2 | Samples | 0.0512 | 2 | B | 0.0545 | 4.87 | 50.85 | TOC Waste Water
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 21:01 | 14 | Clean | Clean |  | 4 | (Clean) |  | -1.02 | 50.101
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 21:01 | 14 | Clean | Clean |  | 4 | (Clean) |  | 6.32 | 50.187
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 21:01 | 14 | Clean | Clean |  | 4 | (Clean) |  | 3.49 | 50.042
mario (mario)[f965748c-c0fa-4399-8f56-1f555046e219] | Fusion | 1 | MGIR250124_CB Sample | 2025-01-24 21:01 | 14 | Clean | Clean |  | 4 | (Clean) |  | 2.26 | 50.846

## Calibration
Sample ID | Curve Value | NDIR (ADJUSTED ABS) | Average | %RPD
[TOC] TOC 20ppm CAL [DI Water] | 0 | 4.46
[TOC] TOC 20ppm CAL [DI Water] | 0 | 4.67 | 4.565 | 4.6002
[TOC] TOC 20ppm CAL [0.4] | 0.4 | 10.44
[TOC] TOC 20ppm CAL [0.4] | 0.4 | 10.41 | 10.425 | 0.28777
[TOC] TOC 20ppm CAL [1] | 1 | 20.31
[TOC] TOC 20ppm CAL [1] | 1 | 19.42 | 19.865 | 4.4802
[TOC] TOC 20ppm CAL [2] | 2 | 35.1
[TOC] TOC 20ppm CAL [2] | 2 | 34.74 | 34.92 | 1.0309
[TOC] TOC 20ppm CAL [5] | 5 | 80.88
[TOC] TOC 20ppm CAL [5] | 5 | 81.77 | 81.325 | 1.0944
[TOC] TOC 20ppm CAL [10] | 10 | 159.12
[TOC] TOC 20ppm CAL [10] | 10 | 162.29 | 160.705 | 1.9726
[TOC] TOC 20ppm CAL [20.000 ppm] | 20 | 298.65
[TOC] TOC 20ppm CAL [20.000 ppm] | 20 | 310.31 | 304.48 | 3.8295
%RPD <10

## QC
Sample ID | PPM C | Mean ppm C | %R | %RPD
MDL | 0.295
MDL | 0.2529 | 0.27395 | 136.975 | 15.3678 | *%R MDL- 45-145
QCS | 17.0337
QCS | 17.995 | 17.5144 | 97.3019 | 5.4886 | *%R 85-115
CCV1 | 9.7991
CCV1 | 9.7815 | 9.7903 | 97.903 | 0.17977
CCV2 | 9.7554
CCV2 | 10.2435 | 9.9994 | 99.9945 | 4.8813
QCB | 0.0553
QCB | 0.0637
CCB1 | 0
CCB1 | 0
CCB2 | 0.048
CCB2 | 0.0545
Average | 0.036917

## Samples
Sample ID | PPM C | Mean ppm C | %RPD | umol/L C
MWR N/S | 53.9766
MWR N/S | 55.7162 | 54.8464 | 3.1718 | 4566.5113

## Reported Results
Sample ID | umol/L C
MWR N/S | 4566.5113

## C.A.L
Attachment 1
CESE
Corrective Action Log
Analyst: | Mike Glinski, Ilana Robbins |  | Instrument: Fusion
Date of Analysis: | 2025-01-24 |  | Method: TOC | Waste Water, 1:10 Dilution
Project Number(s)/Batch Numbers | Excursions | Affected Samples | Criteria Comparison | Corrective Action/Explanation
Below Detection Limit: Retest
Above Curve: Dilute and Retest
Reviewed by: | Mike Glinski |  |  | 2025-01-28
| QA/QC Reviewer |  |  | Date
| Project Investigator |  |  | Date